## Sheet1
Grouping | Item# | Description | Manufacturer | Model Number | Serial No | Description and use | Equipment limits including ranges | Accuracy | Frequency of calibration
environmental | 9 | Temp Humidity Chamber | Thermotron | 2800 | 31997 | conditioning and stressing samples for specific tests | -87C to +190.6C. 5 to 100% R.H | ±2°C and ±1 % R.H. | 1 year
environmental | 18 | Perkin Elmer | Perkin Elmer | DMA/7E |  | used for TG, ΔTg and T260/T288 testing | RT to 450° | ± 1°C | 6 months
environmental | 29 | ISO Temp Oven | Precision Oven | 25EG | 601031161 | preconditoning of samples prior to testing | RT to 200°C | ±2°C | 1 year
hand tool | 2 | Digital Caliper Mitutoyo | Mitutoyo | CD-6-CS | 218107 | dimensional measurement of samples | 0-150 mm or 0-6 inches | 0.0005" | 2 years
hand tool | 5 | Micrometer Mitutoyo | Mitutoyo |  | 193-211 | dimensional measurement of samples | 0-1" | ±0.0001" | 1 year
hand tool | 7 | Scale Denver Instument Co | Denver Instuments | TR-203 | T0116255 | accratley weighing lab solutions including etchants and potting compounds | 0 to 210 grams | ± of applied load | 1 year
hand tool | 8 | Peel Tester | Shimpo | MF Shimpo | 22159 | performing copper peel testing of laminate | 0-2 punds force | ±0.1 lbf | 2 years
hand tool | 10 | Digital Height Gauge | Photocircuits | D5VA |  | dimensional measurement of samples | 0 to 230 mils | ±0.005" | 1 year
hand tool | 11 | Digital Caliper Mitutoyo | Mitutoyo | CD-6-CS | 237183 | dimensional measurement of samples | 0-150 mm or 0-6 inches | ±0.0005" | 2 years
hand tool | 22 | Pin Gage | Meyer | M-1 .061/.25 |  | hole diameter measurement | 0.61 to 0.132" | ±0.005" | 1 year
hand tool | 23 | Pin Gage | Meyer | .22/1.50 | M0MM501 | hole diameter measurement | 0.22 to 1.50 mm | ± 0.01 mm | 1 year
hand tool | 24 | Pin gage | Meyer | .011/.060 | M0*06/14 | hole diameter measurement | 0.11 to 0.60 " | ±0.005" | 1 year
hand tool | 38 | digital thermometer | Fluk Mfg | CNXt3000 | 27030032 | temperature measurement | neg 200°to 1372°C | ±(0.5%+0.3°C) counts | 1 year
meter | 1 | MegohmMeter / IR Tester QuadTech 1865 | QuadTech | 1865 | 298951 | meaurement of insulation resistance of test coupons | 1E+4 ohms to 3.96 E+15 ohms | resistance value dependent (available upon request) | 1 year
meter | 3 | Multimeter | Fluk Mfg | 25 multimeter | 5715027 | measurement of volts and continuity | 0 to 1000 volts/ 0- 32Mohms | resistance value dependent (available upon request) | 1 year
meter | 20 | LCR Meter | Gwinstek | EL 190622 | EL190622 | measurement of capacitance and resistance | 0.00001pF to 999µF. 0.00001Ω to 999KΩ | ± 0.5% | 1 year
meter | 21 | OHM Meter | Keithly | 580 micro ohm meter | 591387 | 4 probe resistance measurement | 200mΩ to 200KΩ | ± 0.5% | 1 year
meter | 26 | Dielectric Strength Tester | Criterion | Criterion Instruments LTD | 6773 | measurement of dielectric break down of laminate | 0 to 2750 AC volts | 10 volts | 1 year
meter | 28 | Fischer Scope XRF | Fischer | XDal 237 | 110007453 | non destructive measurement of plating thickness and composition | 0.3 mils to 2 mils | ± 5% | 1 year
ref std | 4 | Starrett Gauge Block Grade 2 | Starrett |  | 90710 | used for low mag micrscope measurement calibration | 0.020 to 12 inches | ASME.B89.1.9 | 2 years
ref std | 16 | Granite Surface Plate | Starrett | 18X24X3 OLB | 70651 | reference plane for measurements | n/a | 0.00015 TIR | 2 year
ref std | 17 | Optical Micrometer | Unitron | 0.001in Unitron | A01101-104 | calibrate inverted microscopes | 0 to 40 mils | ±0.0002" | 1 year
ref std | 25 | Troemner Weight | Troemner |  | 8393 | calibration of the wetting balances | 100mg to 4 grams | ± 1% | 2 years
scope | 12 | Zeiss Inverted microscope | Zeiss | Axiovert 25 | 666103 | inspection of previously potted cross sections | 50X to 1000X | 0.15 thou @25X to 0.01 thou at 1000X | 6 months
scope | 13 | Measurment System Buehler Omnimet | Buehler | Omnimet |  | image capture and mesurement software | 10 microinches to 150 mils | 0.15 thou @25X to 0.01 thou at 1000X | 6 months
scope | 14 | Nikon Stereo Scope | Nikon | SMZ800 | 540422 | inspection of non potted samples | 0.63X to 63X (with option for 1.26X to 126X) | 0.47 thou@ 10X to 0.09 thou at 63X | 6 months
scope | 15 | Zeiss Inverted microscope | Zeiss | Axi Observer A1 | 3832001180 | inspection of previously potted cross sections | 25X to 1000X | 0.15 thou @25X to 0.01 thou at 1000X | 6 months
scope | 27 | Nikon Stereo Scope | Nikon | SMZ 1000 | 8e+198 | inspection of non potted samples | 8X to 80X (with option for 16X to 160X) | 0.15 thou @ 8X to 0.06 thou at 80X | 6 months
solderability testing | 6 | Wetting Balance Menisco ST50 | Metronelec | ST50 | 521 | wetting force and angle measurement of solderable surfaces and devices | -50 to +50 mN/mm | 0.01mN | 1 year
solderability testing | 19 | Wetting Balance Menisco ST 88 | Metronelec | ST 88 | ST-88 | wetting force and angle measurement of solderable surfaces and devices | -50 to +50 mN/mm | 0.01 mN | 1 year
scope | 92 | Olympus Confocal microscope | Olympus | Lext |  | non contact surface morphology measurement |  |  | 1 year
solderability testing | 31 | SAT Steam Age Controller | Mountaingate Engineering Inc | µSAT | 89020-1-01 | preconditoning of samples prior to testing | RT to 100°C | ±2°C | 1 year
scope | 93 | Zeiss Upright scope | Zeiss | Axotron 2 |  | inspection of previously potted cross sections | 25X to 2000X | 0.15 thou @25X to 0.01 thou at 1000X | 1 year
scope | 106 | Nikon stereo scope | Nikon | SMZ-U |  | inprocess inspection of cross sections | 7.5X to 75X | no measurement capability | 1 year
scope |  | scanning electron microscope | Hitachi | 2460N |  | failure analysis | 40X to 200000X | magnification dependent
scope |  | scanning electron microscope | Zeiss | EVO50 |  | failure analysis | 18X to 1000000X | magnification dependent | 1 year
scope |  | EDS | Bruker | X-flash |  | failure analysis | elements > Boron | element dependent | each use
grinder |  | semi automated cross section grinder | allied | Metprep4 PH6 |  | cross section preparation | 1 1/8 and 1 1/2" mounts | no measurement capability | n/a
polisher |  | semi automated cross section grinder | allied | Metprep4 PH6 |  | cross section preparation | 1 1/8 and 1 1/2" mounts | no measurement capability
grinder/ polisher |  | precison decapsulation equipment | allied | multiprep |  | planar decpapsulation of devices | all current electronic devices | 1 micron